- Test parameter default values

In General Screen

Test ID 			= system generated ID

Full Name 		=

Unit 			=

Reagent Blank		= Disabled

Sample Blank		= Disabled

Bichromatic 		= Disabled

In settings screen

Method 		=End Point

Primary filter		=340

Secondary filter		=Disabled

Temperature		=37

Aspiration volume	=500 uL

Cuvette type		=flow cell

T. Factor		= 1.000

Lag time(s)		= 5

Read time		= 5

In Limit screen

R Blank 			=  0.000 to 0.000

Abs Limit  		= 0.000 to 2.300

Ref range		= 0.000 to 0.000

Linearity range		= 0.000 to 0.000

Max Del A/min		= 0.000

In Calibration screen

Test 			= same as full name

Method 		= Factor or 1point

K factor 		= 1.0

No of standards 	= 0

In standard screen

No of standards conc and abs box enabled based on of standard selected by user in calibration screen.

If no of standards selected to 1 then S0 `s default value is water.

Notes:

Flow cell Temperature options   = 37, 30, 25, RT (Room temperature).

- RUN Test procedure based on sample blank on/off and reagent blank of/off.

- Reagent blank – Disabled.

Sample blank – Disabled.

There is no reagent blank and reagent test.

System will Directly go to sample test (No sample blank test).

Calculation to find the concentration: (Example equation for end point).

Conc = Abs of sample X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = (Conc of standard/ Abs of standard).

- Reagent blank - Enabled.

Sample blank – Disabled.

reagent blank test needs to be done. (No reagent test).

sample test needs to be done. (No sample blank test).

Calculation to find the concentration: (Example equation for end point).

Conc = (Abs of sample – Abs Reagent blank) X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = ((Conc of standard)/ (Abs of standard- abs reagent blank)).

- Reagent blank - Disabled.

Sample blank – Enabled.

There is no reagent blank and reagent test.

Sample blank and sample tests need to be done.

Calculation to find the concentration: (Example equation for end point).

Conc = (Abs of sample – Abs sample blank) X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = ((Conc of standard)/ (Abs of standard- abs standard blank)).

- Reagent blank - Enabled.

Sample blank – Enabled.

Reagent blank and reagent tests need to be done.

Sample blank and sample tests need to be done.

Calculation to find the concentration: (Example equation for end point).

Conc = ((Abs of sample – Abs sample blank) – (Abs Reagent - Abs Reagent blank)) X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = ((Conc of standard)/

(Abs of standard- abs standard blank) – ((Abs Reagent - Abs Reagent blank)).

For kinetics and fixed time calculations.

Kinetics  :

Conc = ((Delta Abs of sample –Delta Abs sample blank) – (Delta Abs Reagent - Delta Abs Reagent blank)) X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = ((Conc of standard)/

(Delta Abs of standard- Delta abs standard blank) – ((Delta Abs Reagent - Delta Abs Reagent blank)).

Fixed time  :

Conc = (((A2-A1) Abs of sample –(A2-A1) Abs sample blank) – ((A2-A1) Abs Reagent - (A2-A1) Abs Reagent blank)) X Factor.

Factor is programmed or calculated from a standard measurement according to

Factor = ((Conc of standard)/

((A2-A1) Abs of standard- (A2-A1) abs standard blank) – (((A2-A1) Abs Reagent - (A2-A1) Abs Reagent blank)).